**Toxoplasma Gondii Ag Rapid Test**

**Cat No.: H81032**

1. **INTENDED USE**

Chemicos™ Toxoplasma Gondii Ag Rapid Test is a sandwich lateral flow immunochromatographic assay for the qualitative detection of Toxoplasma gondii antigen (Toxo Ag) in animal’s feces or serum.

**Assay Time:** 5-10 min

**Sample:** Feces or serum.

1. **PRINCIPLE OF THE ASSAY**

Chemicos™ Toxoplasma Gondii Ag Rapid Test is based on sandwich lateral flow immunochromatographic assay. The test device has a testing window. The testing window has an invisible T (test) zone and C (control) zone. When sample is applied into the sample hole on the device, the liquid will laterally flow on the surface of the test strip. If there is enough Toxoplasma gondii antigen in the sample, a visible T band will appear. The C band should always appear after a sample is applied, indicating a valid result. By this means, the device can accurately indicate the presence of Toxoplasma gondii antigen in the sample.

1. **KIT COMPONENT**
* 20×foil pouches, each containing one cassette and a desiccant
* 20×assay buffer
* 20×swab sticks
* Product Manual
1. **TEST PROCEDURE**

****

* Collect animal’s fresh feces with a swab stick.
* Insert the wet swab into the provided assay buffer tube. Agitate it to assure good sample extraction.
* If using serum sample, drop 1 drop into the provided assay buffer and mix well.
* Take out the cassette from the foil pouch and place it horizontally.
* Gradually drip 3 drops of sample extraction into the sample hole.
* Interpret the result in 5-10 minutes. Result after 10 minutes is considered as reference.
1. **INTERPRETATION OF RESULTS**

**Positive**: The presence of both C band and T band, no matter T band is clear or vague.

**Negative:** Only clear C band appears.

**Invalid:** No colored band appears in C zone, no matter whether T band appears.

****

1. **STORAGE**

The kit can be stored at room temperature (2-30°C). The test kit is stable through the expiration date (24 months) marked on the foil pouch. **DO NOT FREEZE**. Do not store the test kit in direct sunlight.

1. **PRECAUTIONS**
* For best results, please strictly adhere to these instructions.
* All reagents must be at room temperature before running the assay.
* Do not remove test cassette from its pouch until immediately before use.
* Do not reuse the test kit.
* Do not use the test beyond its expiration date marked on the foil pouch.
* The components in this kit have been quality control tested as standard batch unit. Do not mix components from different lot numbers.
1. **LIMITATION**

Chemicos™ Toxoplasma Gondii Ag Rapid Test is for in vitro veterinary diagnosis use only. All results should be considered with other clinical information available from veterinarian. For an accurate result, It is suggested to apply other methods such as PCR for final determination in practice.